# FICHE  META_INFORMATION_PARAMETRES

# (à remplir par le responsable du paramètre)

### Nom du DATASET / Data SET NAME

CDOM absorption

### PROJET-ETUDE / PROJECT TITLE

Campaign  NAME :   	GreenEdge – ICECAMP – 2015	LEG :

Date begin : 2015-142 (Julian day)

Date end : 2015-191 (Julian day)

Chief Scientist: Marcel Babin

Address :

Université Laval

UMI Takuvik

1045 avenue de la médecine

Québec, QC, G1V0A6, Canada

### OPERATION (if Relevant)

Sampling method : CTD, Ice core

Station number-Cast number :  142-191 (Julian day)

Operation code :

### RESPONSABLE SCIENTIFIQUE du paramètre / PI of the parameter

| # Nom / name | # adresse / # address | # téléphone / phone number | # fax / # fax number | # adresse mél / # email address |
| --- | --- | --- | --- | --- |
| # A. Mastsuoka | # Takuvik/UL | # +1 418 656 2131 #3198 | # +1 418 656 2339 | # Atsushi.Matsuoka@takuvik.ulaval.ca |

### DATASET contact

| # Nom / name | # adresse / # address | # téléphone / phone number | # fax / # fax number | # adresse mél / # email address |
| --- | --- | --- | --- | --- |
| # J. Ferland | # Takuvik/UL | # +1 418 656 2131 #8831 | # +418 656 2339 | # joannie.ferland@takuvik.ulaval.ca |

### INFORMATION GEOGRAPHIQUES /  GEOGRAPHIC INFORMATION

Predefined site (if relevant):

Location: Polar heaven ice camp site

LATITUDE: ???

LONGITUDE: ???

### DESCRIPTION DES INSTRUMENTS /  INSTRUMENTS DESCRIPTION (if Relevant)

Instrument Type: Liquid waveguide, UltraPath

Manufacturer: World Precision Instruments Inc.

Model:

Instrument Features / Calibration:

### DESCRIPTION DES PARAMETRES /  PARAMETERS DESCRIPTION

# Ce qui a été collecté, mesuré et comment / How was the parameter collected and measured (include references for analytical methods)?

Sampling: Water column, bottom of the ice core (low and high snow conditions)

Analytical procedure : (briefly, could be a short recall to a published reference):

The detailed methodology to determine light absorbance of CDOM using an UltraPath (World Precision Instruments, Inc.) is documented in Bricaud et al. [2010] and Matsuoka et al. [2012]. Briefly, a sample was collected from a CTD/Niskin bottle (or from a plastic container when using the barge or zodiac) into pre-rinsed glass bottles covered with aluminium foil. Those samples were filtered immediately after sampling, in dim light, using 0.2 μm GHP filters (Acrodisc Inc.) pre-rinsed with 200 ml of Milli-Q water. Filtered samples were then pumped into the sample cell of the Ultrapath instrument using a peristaltic pump. Absorbance spectra were measured from 200 to 735 nm with 1 nm increments with reference to a salt solution (the salinity of the reference was adjusted to that of the sample, ± 2 salinity unit), prepared with Milli-Q water and granular NaCl precombusted in an oven (at 450 ˚C for 4 hours). The sample tube was cleaned between measurements with diluted solutions of detergent, high reagent grade MeOH, 0.1N HCl, and with Milli-Q water. The cleanliness of the tube was checked by measuring the absorption spectrum of the reference water between each sample.

Temperature differences between reference and sample were minimized as far as possible, but could not be always avoided. As absorbance of pure water depends on temperature, this sometimes resulted in an underestimate of absorbance mostly beyond 700 nm for some samples. No temperature correction was applied to absorbance spectra in the data. To minimize the temperature and salinity effects on the absorbance among samples, the averages of the measured values of ODCDOM(λ) over a 5-nm interval around 685 nm (ODnull,CDOM) was assumed to be 0 and the ODCDOM(λ) spectrum was shifted accordingly [Pegau et al., 1997; Babin et al., 2003]:

aCDOM(λ)=2.303[ODCDOM(λ)−ODnull, CDOM]l,			(Eq. 1)

where 2.303 is a factor for converting base e to base 10 logarithms, and l is the optical pathlength (m). In most cases, a 2 m optical pathlength was used for the measurement, except for coastal waters at the Mackenzie River mouth where a 0.1 m optical pathlength was used. The presence of microbubbles in the sample tube was minimized by using a peristaltic pump. When not totally removed, microbubbles induced a significant (and artificial) absorption detectable in the infrared, and the corresponding absorption spectra were discarded.

Spectra of aCDOM(λ) can be expressed as exponential functions as follows [e.g., Bricaud et al., 1981; Babin et al., 2003]:

aCDOM(λ)=aCDOM(λr)e(−SCDOM(λ−λr)),				(Eq. 2)

where λr is a reference wavelength (440 nm in this study), and SCDOM denotes the spectral slope of CDOM. SCDOM was calculated by fitting a non-linear regression to the data from 350 to 500 nm [Babin et al., 2003].

Units: m-1

Sensor Precision: TBD

# Décrire quels types de données sont nécessaires pour vous compléter votre propre jeu de données avant envoi à la base de données, et estimer le délai avant la disponibilité de vos données pour la base de données / Post-cruise data analysis/treatment required, and the time frame for this

Estimated Date of Delivery : TBD

### REFERENCES BIBLIOGRAPHIQUES

Babin, M., Stramski, D., Ferrari, G. M., Claustre, H., Bricaud, A., Obolensky, G., and 　Hoepffner, N.: Variations in the light absorption coefficients of phytoplankton, nonalgal particles, and dissolved organic matter in coastal waters around Europe, J. Geophys. Res., 108, doi:10.1029/2001JC00082, 2003.

Bricaud, A., Babin, M., Claustre, H., Ras, J., and Tièche, F.: Light absorption properties and absorption budget of Southeast Pacific waters, J. Geophys. Res., 115, doi:10.1029/2009JC005517, 2010.

Matsuoka, A., Bricaud, A., Benner, R., Para, J. Sempere, R., Prieur, L., Belanger, S., and Babin, M.: Tracing the transport of colored dissolved organic matter in water masses of the Southern Beaufort Sea: relationship with hydrographic characteristics, Biogeosciences, 9, doi: 10.5194/bg-9-925-2012, 2012.