META_INFORMATION_PARAMETRES / PARAMETERS

(à remplir par le responsable du paramètre /    to be filled by the PI    )

| Date de création | 29th June 2010 | created on |
| --- | --- | --- |
| Date: de dernière mise à jour | 29th June 2010 | last update |

| ### PARAMETRES CONCERNES | ### PARAMETERS |
| --- | --- |

| Mode d’échantillonnage | No parametre | nom parametre | Status : [non] valid |
| --- | --- | --- | --- |
| Sampling | parameter number | name of the parameter | Status : [not] valid |
| Continuous sampling above the surface | 276 | bbp(470nm) | Valid, restricted |
| Continuous sampling above the surface | 276 | bp(526nm) | Valid, restricted |

NB : The third wavelength did not work

| ### PROJET ETUDE | ### PROJECT TITLE |
| --- | --- |

BOUM

| ### RESPONSABLE SCIENTIFIQUE | ### PRINCIPAL INVESTIGATOR |
| --- | --- |

| # Nom / name | # adresse / # address | # téléphone / phone number | # fax / # fax number | # adresse mél / # email address |
| --- | --- | --- | --- | --- |
| Michael | Behrenfeld | +1-541-737-5289 | +1-541-737-3573 | mjb@science.oregonstate.edu |
| Giorgio | Dall’Olmo | +1-541-737-5274 | +1-541-737-3573 | giorgiod@science.oregonstate.edu |

| ### BREVE DESCRIPTION DU PROJET | ### BRIEF DESCRIPTION OF PROJECT |
| --- | --- |

| ### DESCRIPTION DES PARAMETRES | ### PARAMETERS DESCRIPTION |
| --- | --- |

# Ce qui a été mesuré et comment  (méthode analytique et/ou références )/

# What did you measure and how did you do it (include references for analytical methods)?

Measurements of particulate backscattering coefficient (bbp) were collected on water continuously pumped from about 9 m below the sea surface.  The instruments used were two 3-wavelength WETLabs ECO-BB3 installed in custom made flow-through chambers.  Only the blue and green channels were used because the red channels showed significant drifts in the calibration coefficients.  To account for the dissolved signals, instrumental drifts, and biofouling, a custom-made automatic valve directed the bulk seawater through a 0.2-um Cole Palmer nylon cartridge filter, for ten minutes every hour.  To compute bbp, the median values of the bbp coefficients measured during each of the 0.2-um filtration times were linearly and piecewise interpolated between hours and subtracted from the bulk measurements.  Particulate backscattering coefficients were calculated as in Dall'Olmo et al., 2009.  The data in the file are the median of the two independently calibrated instruments.

# Stratégie d'échantillonnage /

# Sampling strategy

Data were recorded continuously at 4Hz and binned every minute.

# Décrire quels types de données sont nécessaires pour vous compléter votre propre jeu de données avant envoi à la base de données, et estimer le délai avant la disponibilité de vos données pour la base de données /

# Post-cruise data analysis/treatment required, and the time frame for this

# Estimations des erreurs, précision, sensibilité des données /

# Error estimates, precision and accuracy of the data

The data presented are the median of two independently calibrated instruments.  The median ratio (± ½ the difference between 84th and 16th percentiles, which is equal to one standard deviation if the distribution is normal) of bbp(470) for the two instruments was 0.920±0.066.  For bbp(526) it was 0.993±0.050.  Thus, we believe these measurements have an accuracy of better than 10%.  The precision of the measurements can be estimated by computing the dispersion of the data over the relevant time scales.

| ### FICHIERS | ### FILES |
| --- | --- |

# Nom de fichier de données /

# File name

Continuous_data_particlulate_backscattering_coefficients.dat

# Explication des têtes de colonne, des unités et des abréviations utilisées dans le fichier de données /

# Data file structure, units, label columns, …

There are 5 columns of data separated by commas.  Each column is labeled (first row), units are reported on row 2.

The following table provides an example:

Time_since_2007/12/31_23:59:59	Latitude	Longitude	bbp(470)	bbp(526)

(Decimal days)	(Decimal degrees North)	(Decimal degrees East)	(1/m)	(1/m)

1.85E+002	3.41E+001	1.84E+001	7.66E-004	7.40E-004

1.85E+002	3.41E+001	1.84E+001	7.87E-004	7.19E-004

1.85E+002	3.41E+001	1.84E+001	7.94E-004	7.49E-004

1.85E+002	3.41E+001	1.84E+001	7.94E-004	7.35E-004

| ### RESULTATS PRELIMINAIRES | ### RESULTS |
| --- | --- |

| ### REFERENCES BIBLIOGRAPHIQUES | ### REFERENCES - PAPERS |
| --- | --- |

Dall’Olmo, G., T.K. Westberry, M.J. Behrenfeld, E. Boss, W.H. Slade (2009). Significant contribution of large particles to optical backscattering in the open ocean.   Biogeosciences, 6, 947–967.

Wayne H. Slade, Emmanuel Boss, Giorgio Dall'Olmo, M. Rois Langner, James Loftin, Michael J. Behrenfeld, Collin Roesler, Toby K. Westberry (in press) Underway and moored methods for improving accuracy in measurement of spectral particulate absorption and attenuation.  Journal of Atmospheric and Oceanic Technologies, (doi: 10.1175/2010JTECHO755.1)