# FICHE  META_INFORMATION_PARAMETRES

# (à remplir par le responsable du paramètre)

### Nom du DATASET / Data SET NAME

Data set Name : Dissolved Organic Phosphorus

### PROJET-ETUDE / PROJECT TITLE

Voir le fichier 	DeWEX_META_INFO_compil.doc

### OPERATION & CAMPAGNE & PIs / OPERATION & CAMPAIGN&PIs

Sampling method : 12 Niskin bottles-rosette

Station number-Cast number :

| DEWEX LEG 1 | DEWEX LEG 2 |
| --- | --- |
| Station number/Cast number | Station number/Cast number |
| 002 – 003 – 005 – 007 – 009 – 011 – 014 – 017 – 018 – 019 – 022 – 023 – 025 – 028 – 031 – 035 – 039 – 043 – 045 – 048 – 052 – 059 – 060 – 063 – 066 – 071 | 001– 006 – 009 – 013 – 016 – 019 – 020 – 022 – 025 – 028 – 031 – 037 – 042 – 045 –048– 051 – 055 –061 –067 – 071 – 073 – 074 – 078 – 081 – 083 – 084 – 085 – 087 – 088 – 089 – 091 – 093 – 095 – 098 - 099 |

Operation code : Dx CTD xxx / Niskin x

### RESPONSABLE SCIENTIFIQUE du paramètre / PI of the parameter

| # Nom / name | # adresse / # address | # téléphone / phone number | # fax / # fax number | # adresse mél / # email address |
| --- | --- | --- | --- | --- |
| # Mireille Pujo-Pay | # LOMIC UMR 7621 avenue du Fontaulé 66650 Banyuls/mer | # 04 68 88 73 51 | # 04 68 88 73 95 | # pujopay@obs-banyuls.fr |
| # Pascal Conan | # LOMIC UMR 7621 avenue du Fontaulé 66650 Banyuls/mer | # 04 68 88 73 52 | # 04 68 88 73 95 | # pascal.conan@obs-banyuls.fr |

### DATASET contact

| # Nom / name | # adresse / # address | # téléphone / phone number | # fax / # fax number | # adresse mél / # email address |
| --- | --- | --- | --- | --- |
| # Mireille Pujo-Pay | # LOMIC UMR 7621 avenue du Fontaulé 66650 Banyuls/mer | # 04 68 88 73 51 | # 04 68 88 73 95 | # pujopay@obs-banyuls.fr |
| # Pascal Conan | # LOMIC UMR 7621 avenue du Fontaulé 66650 Banyuls/mer | # 04 68 88 73 52 | # 04 68 88 73 95 | # pascal.conan@obs-banyuls.fr |

### INFORMATION GEOGRAPHIQUES /  GEOGRAPHIC INFORMATION

Voir le fichier 	DeWEX_META_INFO_compil.doc

### DESCRIPTION DES INSTRUMENTS /  INSTRUMENTS DESCRIPTION

Instrument Type: Autoanalyseur

Manufacturer: Bran-Luebbe - SEAL

Model: AA-III HR

Instrument Features / Calibration: The slope of a calibration curve carried out with artificial seawater (salinity 38). Deionised water is used for the baseline as reference zero nutrient.

### DESCRIPTION DES PARAMETRES /  PARAMETERS DESCRIPTION

# Ce qui a été collecté, mesuré et comment / How was the parameter collected and measured (include references for analytical methods)?

Sampling: Samples were collected from Niskin bottles in 500 ml calcinated Schott glass bottles. The samples were then filtered through 2 combusted GF/F filters. 20ml of the filtered samples were transferred to 20 ml teflon vials and frozen at -20°C immediately.

Analytical procedure : (briefly, could be a short recall to a published reference):

Following the addition of 2.5 ml of the oxidative reagent (boric acid + sodium hydroxide + potassium peroxodisulfate), samples were then heated at 120°C during 30 min. After cooling, the concentration of total dissolved phosphorus was determined using continuous flow analysis (AAIII HR): the absorbance of the blue complex formed with the samples and the reagents is measured at 820nm. Dissolved Organic Phosphorus is obtained by difference between Total Dissolved Phosphorus and Dissolved Inorganic Phosphorus.

Units:  µmol.L-1

Sensor Precision:

Precision : The precision 5%, was the relative standard deviation obtained from 10 determinations (oxidation and colorimetric detection) of the same sample

Limite de détection : 0.01µmol/l

Validation des données obtenues : Accuracy was determined by measuring certified standard CERTIPUR/Merck. For a concentration of 1.05µmol/l the accuracy was 2 %

# Décrire quels types de données sont nécessaires pour vous compléter votre propre jeu de données avant envoi à la base de données, et estimer le délai avant la disponibilité de vos données pour la base de données / Post-cruise data analysis/treatment required, and the time frame for this

Estimated Date of Delivery : October 2013

### REFERENCES BIBLIOGRAPHIQUES

Alain Aminot, Roger Kérouel : Dosage automatique des nutriments dans les eaux marines : méthodes en flux continu. Ed. Ifremer. 2007

Murphy, J. and Riley, J.P., 1962.

A modified single solution method for the determination of phosphate in natural waters .

Anal.Chim.Acta 27, 31-36.

Pujo-Pay, M., Conan, P., and Raimbault, P.: Particulate and dissolved organic nitrogen and phosphorus in the north western Mediterranean Sea (EROS 2000 - Discovery cruise 1993), edited by: Martin, J. M., and Barth, H., 79-86, 1995.

Pujo-Pay, M., Conan, P., and Raimbault, P.: Excretion of dissolved organic nitrogen by phytoplankton assessed by wet oxidation and N-15 tracer procedures, Mar. Ecol. Prog. Ser., 153, 99-111, 1997