# FICHE  META_INFORMATION_PARAMETRES

# (à remplir par le responsable du paramètre) / (to be filled by the person in charge of the parameter)

### PARAMETRES CONCERNES / Parameters

Name of the Parameter : Phytoplankton Abundance by Flow cytometry

### OPERATION & CAMPAGNE & PIs / OPERATION & CAMPAIGN&PIs

Sampling method : Liquid samples were collected from the Rosette

Station number-Cast number :  A3-1, TNS-10, TNS-9, TNS-8, TNS-7, TNS-6, TNS-5,      TNS-4, TNS-3, TNS-2, TNS-1, R2, E-1, TEW-1, TEW-2, TEW-3, TEW-4, E-2, TEW-5, TEW-7, TEW-8, E-3, F-L, E-4, A3-2, E-5.

Operation code : CTD-004, CTD-006, CTD-007, CTD-008, CTD-009, CTD-010, CTD-011, CTD-012, CTD-013, CTD-014, CTD-015, CTD-017, CTD-027, CTD-035, CTD-037, CTD-038, CTD-042, CTD-043, CTD-044, CTD-046, CTD-047, CTD-050, CTD-063, CTD-081, CTD-094, CTD-107, CTD-114.

### PROJET  ETUDE / PROJECT TITLE

Campaign  NAME : KEOPS2

Projetc Leader: Stéphane Blain (stephane.blain@obs-banyuls.fr)

Address :	Laboratoire d'Océanographie Microbienne (LOMIC)

UMR 7621 CNRS UPMC

avenue du Fontaulé

66650 Banyuls sur mer

Chief Scientist:  Bernard Queguiner (bernard.queguiner@univ-amu.fr)

Address :	Institut Méditerranéen d'Océanographie

Université d'Aix-Marseille

UMR CNRS 7294/UR IRD 235

Campus de Luminy, case 901

F-13288 Marseille Cedex 09

### RESPONSABLE SCIENTIFIQUE du paramètre / PI of the parameter

Name : Ingrid Obernosterer/Nicole Batailler/Philippe Catala

Adresse : Laboratoire d’Oceanographie Microbienne (LOMIC) Observatoire océanologique de Banyuls sur Mer Avenue du Fontaulé 66650 Banyuls sur Mer, France

# Tel : 0468887353

# Fax: 04 68 88 73 95

e-mail : ingrid.obernosterer@obs-banyuls.fr

### DESCRIPTION DES PARAMETRES /  PARAMETERS DESCRIPTION

# Ce qui a été collecté, mesuré et comment / How was the parameter collected and measured (include references for analytical methods)?

Sampling:  5 ml of sea water, fixed with glutaraldehyde (1% final concentration)

Analytical procedure : (briefly, could be a short recall to a published reference):

Seawater samples (collected from the CTD-bottles) were fixed with glutaraldehyde (1 % final concentration). After at least 15 minutes on the bench at room temperature in the dark, samples were frozen in liquid nitrogen then stored at - 80°C until their analysis in the laboratory.

Counts were performed with the FacsCalibur flow cytometer (Becton Dickinson) equipped with an air-cooled argon laser (488 nm, 15 mW). Samples were thawed at room temperature.

Phytoplanktonic cells were enumerated according to SSC, orange (585/42 nm) and red (> 650 nm) fluorescences corresponding to phycoerythrin and chlorophyll pigments, respectively.

The speed of the analysis depends on the cell abundance, typically the volume analyzed was 300 µl. The phytoplankton cell abundance was determined from the flow rate, calculated as the difference between the weight of milli-Q water before and after a 5 min run of the cytometer. The flow rate was determined after every tenth sample tube, and it was stable over the time of analyses. Fluorescent beads (1.002 µm; Polysciences Europe) were systematically added to each analyzed sample as internal standard.

Units:  cells per ml

# Décrire quels types de données sont nécessaires pour vous compléter votre propre jeu de données avant envoi à la base de données, et estimer le délai avant la disponibilité de vos données pour la base de données / Post-cruise data analysis/treatment required, and the time frame for this

Estimated Date of Delivery :  December 2012

# Estimations des erreurs, précision, sensibilité des données / Error estimates, precision and accuracy of the data

One samples per depth was run for the KEOPS2 cruise. From previous experience, the coefficient of variation for this analyses is less than 5%.

### REFERENCES BIBLIOGRAPHIQUES

Obernosterer,I., P. Catala, R. Lami, J. Caparros, J. Ras, A. Bricaud, C. Dupuy, F. van Wambeke and P. Lebaron (2008) Biochemical characteristics and bacterial community structure of the sea surface microlayer in the South Pacific Ocean. Biogeosciences 5: 693-705.