# FICHE  META_INFORMATION_PARAMETRES

# (à remplir par le responsable du paramètre)

### Nom du DATASET / Data SET NAME

Data set Name : Silicate ou Acide Silicique, PO4, NO2, NO2+NO3

### PROJET-ETUDE / PROJECT TITLE

STRASSE

### OPERATION & CAMPAGNE & PIs / OPERATION & CAMPAIGN&PIs

Sampling method : 24 Niskin bottles-rosette

### RESPONSABLE SCIENTIFIQUE du paramètre / PI of the parameter

| # Nom / name | # adresse / # address | # téléphone / phone number | # fax / # fax number | # adresse mél / # email address |
| --- | --- | --- | --- | --- |
| # Pujo Pay Mireille | # LOMIC UMR 7621 avenue du Fontaulé 66650 Banyuls/mer | - +33 4 68 88 73 51 |  | # pujopay@obs-banyuls.fr |

### DATASET contact

| # Nom / name | # adresse / # address | # téléphone / phone number | # fax / # fax number | # adresse mél / # email address |
| --- | --- | --- | --- | --- |
| # Pujo Pay Mireille | # LOMIC UMR 7621 avenue du Fontaulé 66650 Banyuls/mer | - +33 4 68 88 73 51 |  | # pujopay@obs-banyuls.fr |

### DESCRIPTION DES INSTRUMENTS /  INSTRUMENTS DESCRIPTION

Instrument Type: Autoanalyseur

Manufacturer: Bran-Luebbe - SEAL

Model: AA-III HR

Instrument Features / Calibration: The slope of a calibration curve carried out with artificial seawater (salinity 38). Deionised water is used for the baseline as reference zero nutrient.

### DESCRIPTION DES PARAMETRES /  PARAMETERS DESCRIPTION

# Ce qui a été collecté, mesuré et comment / How was the parameter collected and measured (include references for analytical methods)?

Sampling: Samples were collected from Niskin bottles in two 20 ml PE scintillation vials. Following rinsing, the vials were filled with seawater and frozen at -20°C immediately.

## 7.1.a	Si(OH)4

Analytical procedure : (briefly, could be a short recall to a published reference):

After defrosting, the absorbance of the blue complex formed with the samples and the reagents is measured at 820nm.

Units:  µmol.L-1

Sensor Precision:

Precision standard deviation < 1%

Limite de détection : 0.05µmol/l

Validation des données obtenues : Accuracy was determined by measuring certified standard CERTIPUR/Merck. For a concentration of 7 µmol/l the accuracy was 1 %

## 7.1.b	PO4

Analytical procedure : (briefly, could be a short recall to a published reference):

The absorbance of the blue complex formed with the samples and the reagents is measured at 820nm.

Units:  µmol.L-1

Sensor Precision:

Precision  standard deviation < 2%

Limite de détection : 0.01µmol/l

Validation des données obtenues : Accuracy was determined by measuring certified standard CERTIPUR/Merck. For a concentration of  1.05µmol/l the accuracy was 2 %

## 7.1.c	NO2

Analytical procedure : (briefly, could be a short recall to a published reference):

After defrosting, the absorbance of the pink complex formed with the samples and the reagents is at 540nm.

Units:  µmol.L-1

Sensor Precision:

Precision standard deviation < 2%

Limite de détection : 0.01µmol/l

Validation des données obtenues : Accuracy was determined by measuring certified standard CERTIPUR/Merck. For a concentration of 1.08 µmol/l the accuracy was 1.9 %

## 7.1.d	NO3

Analytical procedure : (briefly, could be a short recall to a published reference):

After defrosting, the nitrate of the samples is first reduced in nitrite by passing through a cadmium-copper column. The absorbance of the pink complex formed with the samples and the reagents is measured at 540nm.

Units:  µmol.L-1

Sensor Precision:

Precision standard deviation < 1%

Limite de détection : 0.02µmol/l

Validation des données obtenues : Accuracy was determined by measuring certified standard CERTIPUR/Merck. For a concentration of 6.5µmol/l the accuracy was 1 %

### REFERENCES BIBLIOGRAPHIQUES

Alain Aminot, Roger Kérouel : Dosage automatique des nutriments dans les eaux marines : méthodes en flux continu. Ed. Ifremer. 2007

Mullin, J.B. and Riley, J.P. (1955) The spectrophotometric determination of silicate-silicon in natural waters with special reference to sea water. Analytica Chim. Acta, 12, 162-70